System Power Controller – Checkout Procedure

S/N _____	Technician __________	Date ____________

Switching Power Supply

___	Apply 5 VDC supply voltage between "7-60V" and "GND" to power the board.

Monitor voltage at "+5V" test pin.

Slowly increase supply voltage to 60 volts and confirm that "+5V" does not exceed 5.5 volts.

___	Change supply voltage down to 12 VDC.

___	Confirm that the 5V POWER LED is lit.

Board Firmware & Configuration

___	Connect PICkit 2 programmer to J11 PROGRAM connector.  Connect programmer to USB port on PC.

Run Microchip MPLAB IDE program:

Project – Open – "System Power Controller.mcp"

Programmer – Connect

Programmer – Program

Disconnect PICkit programmer from J11 PROGRAM connector.

___	Connect PC serial port to J0 DB9 COMM connector.

Run Terra Term Pro and Setup – Serial Port – Baud rate: 9600

Press ENTER key and confirm "SPC>" prompt.

___	To initialize the EEPROM values, enter the following commands:

However, do not initialize A/D calibration values (EWC commands) if board has aleady been calibrated and is not going to be recalibrated.

EWSn	 where n is the board's serial number

EWP0

EWT3840

EWR60

EWC1M34202	 A/D calibration values

EWC1Z0

EWC2M34202

EWC2Z0

EWC3M8346

EWC3Z0

EWB

EWI12345

RESET

___	Enter the "E" command to confirm that all the EEPROM values are correctly set.

___	If not already done, use a permanent marker to write the serial number on the board.

S/N _____	Technician __________	Date ____________

Power Channels

___	Enter the "P" command to check the state of all the power channels.

___	Enter the "P1H" command to turn channel 1 on.  Confirm that the P1 LED is lit.

Enter the "P1L" command to turn channel 1 off.  Confirm that the P1 LED goes out.

___	Repeat for channels 2-9.

___	Enter the "P0H" command to turn on the watchdog relay on channel 0.  Confirm that the relay clicks.

Enter the "P0L" command to turn off the watchdog relay.  Confirm that the relay clicks.

___	Power the board through VOM amp connectors to measure the board's current consumption.

Disconnect the serial connector to J0 DB9 COMM.

Current at supply voltage of 7.0 volts:	__________ mA  ( < 25 mA )

Current at supply voltage of 60.0 volts:	__________ mA  ( < 10 mA )

___	Reconnect the serial connector to J0 DB9 COMM.

Use the "PnH" commands to turn on all channels 0-9, including the watchdog relay.

Current at supply voltage of 7.0 volts:	__________ mA  ( < 130 mA )

Current at supply voltage of 60.0 volts:	__________ mA  ( < 25 mA )

___	Power off and disconnect VOM.

___	Use a 30 ohm, 120 watt load to confirm each power switching channel.

Set the supply voltage to 60.0 volts.

Turn off power supply between each test setup.

___		Watchdog Relay:

Connect "WDI" to 60 volt supply.

Connect "WDO" through 30 ohm load to ground.

Confirm 60 volts at "WDO".

Enter "P0H" command and confirm 0 volts at "WDO".

Enter "P0L" command and confirm 60 volts at "WDO".

___		Power Channels:

Connect "P1-IN" to 60 volt supply.

Connect "P1-OUT" through 30 ohm load to ground.

Confirm 0 volts at "P1-OUT".

Enter "P1H" command and confirm 60 volts at "P1-OUT".

Enter "P1L" command and confirm 0 volts at "P1-OUT".

___		Repeat for power channels 2-8.

___	Channel P9:

Set power supply to 12.0 volts.

Connect "P9" through 120 ohm resistor to 12 volt supply.

Confirm 12 volts at "P9".

Enter "P9H" command and confirm <0.2 volts at "P9".

Enter "P9L" command and confirm 12 volts at "P9".

S/N _____	Technician __________	Date ____________

Watchdog Timer

___	Enter "W10" command to set a 10 second watchdog timeout.  Confirm the following response:

SPC>W10

10  [time=3840, reset=60]

Repeatedly enter the "W" command to follow the watchdog timeout progress.

After 10 seconds the watchdog relay should power on.

The "R" in the response to the "W" command indicates the relay on timeout:

SPC>W

R60  [time=3840, reset=60]

After 60 seconds the watchdog relay should power off again:

SPC>W

3815  [time=3840, reset=60]

Voltage Reference

___	Measure the voltage on pin 5 of  the U4 PIC16F886 semiconductor chip.

A/D Voltage Reference:	__________ volts  ( 2.994 to 3.006 )

Ground Connections

___	With board powered off, check continuity between test point GND and all GND terminals.  All resistances should be < 0.1 ohms.

S/N _____	Technician __________	Date ____________

A/D Converter Operation & Calibration – V1 Channel

Use a calibrated VOM to measure the voltage directly at the V1 terminals.

New Calibration

Enter the "E" command to confirm that the V1 calibration is set to the default values:

Cal V1 Mult   = 34202

Cal V1 Offset = 0

Use the "V1" command to read the A/D output with the following input voltages.

Take an average of 10 readings for each A/D output:

Setting	VOM Reading (y)	A/D Output (x)

10 volts	_______________	_______________

20 volts	_______________	_______________

40 volts	_______________	_______________

60 volts	_______________	_______________

Calculate a linear regression (y = mx + b) and record the values here:

Slope (m)	____________________	x 34202 =	_______________ (Cal V1 Mult)

Offset (b)	____________________	x 1000 =	_______________ (Cal V1 Offset)

Correlation	____________________

Use the "EWC1Mn" command to enter the "Cal V1 Mult" value.

Use the "EWC1Zn" command to enter the "Cal V1 Offset" value.

Enter "RESET" command to apply changes.

Enter "E" command to confirm correct values stored.

Check Calibration

Repeat the test to confirm that the calibration values are working:

Setting	VOM Reading (y)	A/D Output (x)	Error (x – y)

0 volts		_______________	_______________

10 volts	_______________	_______________	_______________

20 volts	_______________	_______________	_______________

40 volts	_______________	_______________	_______________

60 volts	_______________	_______________	_______________

S/N _____	Technician __________	Date ____________

A/D Converter Operation & Calibration – V2 Channel

Use a calibrated VOM to measure the voltage directly at the V2 terminals.

New Calibration

Enter the "E" command to confirm that the V1 calibration is set to the default values:

Cal V2 Mult   = 34202

Cal V2 Offset = 0

Use the "V2" command to read the A/D output with the following input voltages.

Take an average of 10 readings for each A/D output:

Setting	VOM Reading (y)	A/D Output (x)

10 volts	_______________	_______________

20 volts	_______________	_______________

40 volts	_______________	_______________

60 volts	_______________	_______________

Calculate a linear regression (y = mx + b) and record the values here:

Slope (m)	____________________	x 34202 =	_______________ (Cal V2 Mult)

Offset (b)	____________________	x 1000 =	_______________ (Cal V2 Offset)

Correlation	____________________

Use the "EWC2Mn" command to enter the "Cal V2 Mult" value.

Use the "EWC2Zn" command to enter the "Cal V2 Offset" value.

Enter "RESET" command to apply changes.

Enter "E" command to confirm correct values stored.

Check Calibration

Repeat the test to confirm that the calibration values are working:

Setting	VOM Reading (y)	A/D Output (x)	Error (x – y)

0 volts		_______________	_______________

10 volts	_______________	_______________	_______________

20 volts	_______________	_______________	_______________

40 volts	_______________	_______________	_______________

60 volts	_______________	_______________	_______________

S/N _____	Technician __________	Date ____________

A/D Converter Operation & Calibration – V3 Channel

Use a calibrated VOM to measure the voltage directly at the V3 terminals.

New Calibration

Enter the "E" command to confirm that the V1 calibration is set to the default values:

Cal V3 Mult   = 8346

Cal V3 Offset = 0

Use the "V3" command to read the A/D output with the following input voltages.

Take an average of 10 readings for each A/D output:

Setting	VOM Reading (y)	A/D Output (x)

7 volts	_______________	_______________

10 volts	_______________	_______________

15 volts	_______________	_______________

20 volts	_______________	_______________

Calculate a linear regression (y = mx + b) and record the values here:

Slope (m)	____________________	x 8346 =	_______________ (Cal V3 Mult)

Offset (b)	____________________	x 1000 =	_______________ (Cal V3 Offset)

Correlation	____________________

Use the "EWC3Mn" command to enter the "Cal V3 Mult" value.

Use the "EWC3Zn" command to enter the "Cal V3 Offset" value.

Enter "RESET" command to apply changes.

Enter "E" command to confirm correct values stored.

Check Calibration

Repeat the test to confirm that the calibration values are working:

Setting	VOM Reading (y)	A/D Output (x)	Error (x – y)

0 volts		_______________	_______________

7 volts	_______________	_______________	_______________

10 volts	_______________	_______________	_______________

15 volts	_______________	_______________	_______________

20 volts	_______________	_______________	_______________

___	Checkout complete.